## Slide 1
In-Silico Color Rendering Simulation for Whole-Slide Imaging Scanners with a Spectral Database of Human Tissues
Jocelyn Liu1, Mary Barcus, MD2, Wei-Chung Cheng, PhD3
1Thomas Jefferson High School for Science and Technology, Alexandria, VA; 2CDRH/OIR/DMD, 3CDRH/OSEL/DIDSR
SPECTRAL DATABASE RESULTS
Clinical Area: Medical Imaging
METHODS
Digital slide phantoms provide a novel conduit for sharing and evaluating color performance of WSI devices. However, this database is limited in both its sampling size and the nature of the sampled tissues. Future endeavors should expand to a larger size and begin to sample tumor tissues, as well as include the usage of other medical stains such as PAS and Masson’s trichome which resolve different features and encompass larger color gamuts. We hope these public domain resources will provide a new means of efficiency and accessibility in remotely gauging WSI devices’ color performance.
CONCLUSION
INTRODUCTION
Remotely assessing the color performance of whole-slide imaging (WSI) scanners is challenging for the Agency due to the lack of a standardized, histologically meaningful color phantom. In the past, Sponsors submitted calibration data or scanned images, which cannot be evaluated properly without the spectral information of the original glass slides. In this study, eight human tissue samples were measured with a multispectral imaging system to create a spectral database. Based on the spectral data, Sponsors can simulate and submit the output images, so the Agency can compare the data with predicate devices to determine substantial equivalence.
FUNDING SOURCE: ORISE
A digital phantom for Sponsors to demonstrate color performance of WSI devices
REGULATORY RELEVANCE
Publically available digital phantoms allow Sponsors to simulate the output images of WSI devices and therefore simplify the required color performance tests.
Glass Slide [1]
Multispectral Imaging System [2]
WSI simulator
Predicate Device
WSI simulator
New Device
Spectral Data [3]
Images from
Predicate Device
Images from
New Device
?
=
(a) Bladder
(b) Brain
(c) Breast
(d) Colon
(e) Kidney
(f) Liver
(g) Lung
(h) Uterine cervix
1: Superficial (umbrella) cells
2: Intermediate urothelial cells
3: Connective tissue
4: Basal urothelial cells
1: Blood vessel, RBC’s
2: Neuron
3: Neuropil
1: Connective tissue
2: Epithelial cells
1: RBC’s
2: Tubule, cytoplasm
3: Nucleus [glomerulus]
4: Cytoplasm [glomerulus]
5: Tubule, nucleus
1: Hepatocyte, nucleus
2:Hepatocyte, cytoplasm
1: Blood vessel, RBC’s
2: Alveolar walls (cytoplasm)
3: Alveolar walls (nuclei)
1: Superficial cells
2: Intermediate cells
3: Basal cells
4: Stroma
1: RBC’s
2: Muscularis mucosa
3: Epithelial cells (glandular)
4. Lamina propria
COLOR RENDERING SIMULATION
2018 OSEL Annual Student Science Poster Exhibit, August 1, 2018
REFERENCES
[1] US Biomax, Inc., Rockville, MD, USA.
[2] Evaluating color performance of whole-slide imaging devices by multispectral-imaging of biological tissues, SPIE 2017.
[3] To be uploaded onto a GitHub repository.
[4] Mills, S. E. (Ed.). (2007). Histology for pathologists (3rd ed.). Philadelphia, PA: Lippincott. Williams & Wilkins.
Downloaded by Sponsor B
510(k)
submitted
Downloaded by Sponsor A
510(k)
submitted
SE or NSE
A simulator was created to demonstrate how to use the spectral database. The simulator consists of eight stages to simulate the image signal processing (ISP) inside a color camera.
The Sponsors can tune the simulation parameters
to generate
the simulated images.
Poor W.B.
Ideal image
Gamma=1.0
Halogen lamp
Images reproduced from the spectral transmittance data of the eight tissue samples. The insets in each image shows spectra of sample cells and color distribution in CIELAB.
Light Source
Glass Slide (Digital Phantom)
Infrared Filter
Color Filter Array
De-Mosaicing
Gamma
White Balancing (W.B.)
Analog-Digital Converter